ALTA - ROBBINS

MANOMETERS - VALVES - AIR PURIFICATION

Rupture Disk Replacement and Maintenance instructions		Bulletin 73115

Restek P/N: 	26303 	26305	26319 	26323 	21396	26315	26315	26302

26320	26318	26322	26324 	26314	26316	26304

Alta Robbins P/N: 		314-RD 	RDT

General information:  For all Alta Robbins products that use a rupture disk.

The nominal rupture disk burst pressure should be at least 40% higher than the system maximum operating pressure.  Rupture disk burst pressures will be reduced by pulsating pressures, elevated temperature, and internal or external corrosion.  In corrosive applications, regular inspections and disk replacement are necessary.  Use of this rupture disk in vacuum or alternating vacuum / pressure applications may cause premature disk failure.  When vented, the rupture unit presents a potential danger caused by noise and escaping pressure, chemicals or fragmented particles.  The complete rupture disk assembly consists of the disk, PTFE washer (or optional PTFE liner) and holder.  Each holder is etched with the nominal burst pressure and each disk is color coded for identification.

To replace the rupture disk follow these instructions:

1) Remove the holder, disk, and washer (or liner) from the valve body.

2) Install a new PTFE washer (or liner) on the process side of the disk.

3) Install a new disk on top of the PTFE washer or PTFE liner.  The color code should be visible to the operator.

Note:  The PTFE washer or liner can only be used one time without replacement.

4) Thread in the holder.  The holder should be torqued per the following table. Lock tight could be used to ensure the holder does not become loose.

| Alta Robbins P/N | Burst (PSI) Pressure | Color Code | Disk Material | Torque in inch pounds |
| --- | --- | --- | --- | --- |
| RD500 | 500 | Blue | Aluminum | 40 |
| RD500-L (PTFE liner) |  |  |  |  |
| RD1000 | 1000 | Red | Monel | 75 |
| RD1000-L (PTFE liner) |  |  |  |  |
| RD1500 | 1500 | Yellow | Monel | 120 |
| RD1500-L (PTFE liner) |  |  |  |  |
| RD1900 | 1900 | Green | 316 SS | 140 |
| RD1900-L (PTFE liner) |  |  |  |  |
| RD1900-M | 1900 | White | Monel | 140 |
| RD1900-M-L (PTFE liner) |  |  |  |  |
| RD2850 | 2850 | Black | 316 SS | 215 |
| RD2850-L (PTFE liner) |  |  |  |  |

110 South 1200 West Lindon, UT 84042   PH: 801-785-1114   FX 801-785-4333   Email: altarob@digis.net